# GW Instek GSP-9300頻譜分析儀2FSK分析功能DEMO

這次的演示是GSP-9300 頻譜分析儀的2FSK信號分析功能。

這DEMO KIT是個ASK及FSK的信號產生器。此DEMO KIT有3種預設頻率，3種波特率及可設定輸出為連續或叢發模式。

架設方式：

- 連接DEMO KIT到GSP-9300前面板的USB接口

- 用SMA RF CABLE連接DEMO KIT輸出端到GSP-9300的輸入端。

儀器操作:

- 首先設定DEMO KIT的參數.

- 按Option Control鍵 > Demo Kit[F4]鍵>Demo Kit[F1]鍵*2次,啟動DEMO KIT.

- 按 Modulation [F2]鍵, 選擇FSK[F2]鍵

- 按Frequency[F3]鍵,選擇433MHz[F2]鍵

- 按Deviation[F4]鍵, 選擇頻率偏移量為50kHz[F2]鍵

- 按Baudrate[F5]鍵, 選擇波特率為9600[F2]鍵

- 按Output Mode[F6]鍵, 選擇 normal mode.

- DEMO KIT的參數設定結束.開始設定2FSK分析功能觀測FSK的信號.

- 按 Frequency 鍵> Center Freq[F1]鍵>4鍵>3鍵>3鍵> MHz[F2]鍵, 設定中心頻率為433MHz

- 按 Span鍵> Span[F1]鍵 >5鍵>0鍵>0鍵> kHz[F3]鍵, 設定頻段為500kHz

- 在頻譜分析儀上可以觀測到2FSK信號.

- 按 Measure鍵> 2FSK[F3]鍵>2FSK [F1]*2次,啟動2FSK分析功能

- 2FSK信號分析可以測量的參數包含載波功率, FSK頻偏,載波頻率 & 載波頻率偏移測量。使用者可以設定判斷條件於頻偏及載波偏移以快速判斷測試結果。

- 按LimitEdit[F5]鍵> Freq. Deviation [F1]鍵> 5鍵>0鍵>kHz[F3]鍵, 設定FSK頻偏為50kHz.

- 按Carr.Offset[F2]鍵>1鍵>0鍵>kHz[F3]鍵, 設定載波頻率偏移為 10kHz.

- 按Return[F7]鍵

- 按Pass/Fail[F6]*2次,可在螢幕上看到Pass/Fail訊息.

- 以上是GSP-9300 2FSK分析功能的介紹.

- 感謝您的收看

# GW Instek GSP-9300 Spectrum Analyzer 2FSK Function DEMO

Now, I am going to demonstrate the 2FSK signal analysis function of GSP-9300 spectrum analyzer.

This demo kit is an ASK and FSK signal generator. In this demo kit, there are three preset frequency, three Baud rates and selectable continuous or burst output mode.

First of all, connect demo kit with the USB port located at the front panel of GSP-9300. Next, connect demo kit output terminal with GSP-9300 input terminal with SMA RF cable.

- To begin the operation, first, input demo kit parameters.

- Press option control key, demo kit (F4) key, demo kit (F1) key twice to activate demo kit

- Press modulation (F2) key and select FSK (F2) key

- Press frequency (F3) key and select 433 MHz (F2) key

- Press deviation (F4) key and select frequency deviation of 50kHz (F2) key

- Press Baud rate (F5) key and select baud rate of 9600 (F2) key

- Press output mode (F6) key and select normal mode

- This is the end of demo kit parameters input. Next, set up the 2FSK analysis function to observe FSK signals.

- Press frequency key, center freq (F1) key, input 4, 3, 3, MHz (F2) key to set center frequency as 433 MHz.

- Press span key, span (F1) key, input 5, 0, 0, kHz (F3) key to set frequency bandwidth as 500kHz.

- Now, 2FSK signals are shown on the screen.

- Press measure key, 2FSK (F3) key, 2FSK (F1) twice to activate 2FSK analysis function.

- 2FSK signal analysis measures parameters including carrier power, FSK frequency deviation, carrier frequency, and carrier frequency offset. Users can set the criterion in frequency deviation and carrier offset for fast test result determination.

- Press limitedit (F5) key, freq deviation (F1) key, input 5,0, kHz (F3) key to set FSK frequency deviation as 50kHz.

- Press carr.offset (F2) key, input 1,0, kHz (F3) key to set carrier offset as 10kHz.

- Press return (F7) key

- Press pass/fail (F6) twice. Pass/Fail indicator is shown on the screen.

- This concludes the demonstration of GSP-9300 2FSK analysis function.

- Thank you very much for your interest in GSP-9300.